## HP
mfr website | mfr name | model name | mfr number | unit cost | product description | amps | volts | watts | phase | hertz | plug_type | emergency_power Required (Y/N) | dedicated_circuit Required (Y/N) | tech_conect Required | btu | dissipation_type | water_cold Required (Y/N) | water_hot Required (Y/N) | drain Required (Y/N) | water_treated (Y/N) | steam Required(Y/N) | vent Required (Y/N) | vacuum Required (Y/N) | ship_weight | weight | depth | height | width | ada compliant (Y/N) | green certification? (Y/N) | antimicrobial coating (Y/N) | Specification Sheet (pdf) | Brochure (pdf) | Manual/IFU (pdf) | Product URL | CAD (dwg) | REVIT (rfa) | Seismic document | Product Image (jpg) | Product Image
https://www.justrite.com/ | Justrite Manufacturing | Blue Polyethylene 24080 (Countertop) | 24080 | $730.00 | Safely store up to two 4-liter bottles of acids or bases in this durable, compact cabinet. Never worry about corrosion—completely metal free polyethylene design offers excellent chemical resistance against spills or damaging vapors. Includes two polyethylene spill trays, use one for inside the cabinet to contain leaks and remove for easy cleaning. Store the extra tray under the cabinet, and remove for use as a handy work top. Hinge door from either side for best worktop utilization—accepts a padlock (not included) for security. Three application-specific labels identify contents to safety store and segregate incompatible chemicals—ACID, BASE, and CORROSIVE.Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. | [19 empty] | 19.5 | 17.125 | 19.75 | 14.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/30321/name/Hold_Two_4-Liter_Bottles%252C_1_Door%252C_Manual_Close%252C_CorrosivesAcids_Plastic_Safety_Cabinet%252C_Countertop%252C_Blue_-_24080.pdf/ |  |  | https://www.justrite.com/cabinet-acid-poly-ctop-blu-24080 |  |  |  | https://www.justrite.com/media/catalog/product/_/2/_2_4_24080-2_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/2/_2_4_24080-2_1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Blue Wood Laminate | 24120 | $932.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices.Safely store highly corrosive liquids such as phenol, nitric acid and sulfuric acid in these wood laminate storage cabinets. Ideal for smaller quantities used at workbench level. Long lasting wood construction along with stainless steel handles and hinges provide exceptional resistance to corrosion from harsh acids. Shelf adjusts on 1-1/4" (32-mm) centers to accommodate a wide variety of containers. Cabinets include adjustable leg levelers, Haz-Alert™ reflective labels, and double key set for safety and security. | [19 empty] | 45.0 | 16 | 18.5 | 24 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29293/name/Holds_Six_2.5-Liter_Bottles%252C_2_Doors%252C_Manual_Close%252C_Wood_Laminate_Corrosives_Safety_Cabinet%252C_Countertop%252C_Blue_-_24120.pdf/ |  |  | https://www.justrite.com/wood-laminate-corrosives-countertop-safety-cabinet-cap-six-2-1-2-liter-bottles-2-doors-blue-24120 |  |  |  | https://www.justrite.com/media/catalog/product/2/4/24120_2_w_hazalert.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/2/4/24120_2_w_hazalert.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | ChemCor Corrosives/Acids Safety (60 gal., 2 Doors) | 8960222 | $4,428.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. The ChemCor® thermoplastic coating on all interior surfaces provides superior protection against corrosion from drips and off-gassing from harsh chemicals. Non-porous surface resists staining and cleans up easily.ChemCor® Lined Acid Safety Cabinets have the same quality features found in Justrite's flammable cabinets — including double-wall, 18-gauge (1mm) welded steel construction, 2" (51mm) liquid-tight containment sump, dual vents, adjustable shelves, three-point stainless steel bullet self-latching doors with U•Loc™ padlockable handle and Haz-Alert™ reflective warning labels. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance. Interior ChemCor® liner ensures complete corrosion resistance and protection. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.Self Close models meet OSHA, NFPA 30, NFPA 1 and the International Fire Code. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 332.0 | 34 | 65 | 34 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29315/name/60_Gallon%252C_2_Shelves%252C_2_Doors%252C_Self_Close%252C_CorrosivesAcids_Safety_Cabinet%252C_ChemCor%25C2%25AE%252C_Blue_-_8960222.pdf/ |  |  | https://www.justrite.com/chemcor-corrosives-acids-safety-cabinet-cap-60-gallons-2-shelves-2-self-close-doors-blue-8960222 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_8960222_sfe-lg_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_8960222_sfe-lg_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Chemcor Lined Countertop Corrosives/Acids Cabinet | 8904022 | $1,424.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Countertop and Compac Cabinets let you customize your work areas and make the most of available space. Whether adding to existing storage space or making chemicals readily available at workstation locations, these smaller-sized cabinets are the solution when space is at a premium.The ChemCor® thermoplastic coating on all interior surfaces provides superior protection against corrosion from drips and off-gassing from harsh chemicals. Seam-free surface eliminates the need for rivets or fasteners which risk corrosion from seep-through chemicals. Non-porous surface resists staining and cleans up easily.ChemCor® Lined Countertop Acid Safety Cabinets have the same quality features found in Justrite's flammable cabinets — including double-wall, 18-gauge (1mm) welded steel construction, 2" (51mm) liquid-tight containment sump, dual vents, adjustable shelves, three-point stainless steel bullet self-latching doors with U•Loc™ padlockable handle and Haz-Alert™ reflective warning labels. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance. Interior ChemCor® liner ensures complete corrosion resistance and protection. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed.Manual close models meet OSHA and NFPA 30. Cabinets come with a ten-year warranty. Read More | [19 empty] | 57.0 | 17 | 22 | 17 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29296/name/4_Gallon%252C_1_Shelf%252C_1_Door%252C_Manual_Close%252C_CorrosivesAcids_Safety_Cabinet%252C_ChemCor%25C2%25AE_Countertop%252C_Blue_-_8904022.pdf/ |  |  | https://www.justrite.com/chemcor-countertop-corrosives-acids-safety-cabinet-cap-4-gallons-1-shelf-1-m-c-door-blue-8904022 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_8904022_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_8904022_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | ChemCor Undercounter 22 gal. | 8923022 | $2,461.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. The ChemCor® thermoplastic coating on all interior surfaces provides superior protection against corrosion from drips and off-gassing from harsh chemicals. Non-porous surface resists staining and cleans up easily.ChemCor® Lined Undercounter Acid Safety Cabinets have the same quality features found in Justrite's flammable cabinets — including double-wall, 18-gauge (1mm) welded steel construction, 2" (51mm) liquid-tight containment sump, dual vents, adjustable shelves, three-point stainless steel bullet self-latching doors with U•Loc™ padlockable handle and Haz-Alert™ reflective warning labels. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance. Interior ChemCor® liner ensures complete corrosion resistance and protection. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed.Manual close models meet OSHA and NFPA 30. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 165.0 | 22 | 35 | 35 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29306/name/22_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_CorrosivesAcids_Safety_Cabinet%252C_ChemCor%25C2%25AE_Under_Counter%252C_Blue_-_8923022.pdf/ |  |  | https://www.justrite.com/chemcor-undercounter-corrosives-acids-safety-cabinet-cap-22-gallons-1-shelf-2-m-c-doors-blue-8923022 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_8923022_sfe-lg_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_8923022_sfe-lg_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Corrosives/Acids Plastic Safety Cabinet, White | 24004 | $3,190.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices.Solid polyethylene white acid cabinets have excellent chemical resistance to handle the most corrosive chemicals. A protective, all-welded bottom sump and a clean, sleek design complement any laboratory settings. Fully welded, liquid-tight 2" (51-mm) sump contains accidental spills and makes clean up a snap. Door accepts a padlock for security. | [19 empty] | 26.0 | 17 | 22 | 17 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29289/name/Holds_Six_2.5-Liter_Bottles%252C_1_Door%252C_Manual_Close%252C_CorrosivesAcids_Plastic_Safety_Cabinet%252C_Countertop%252C_White_-_24004.pdf/ |  |  | https://www.justrite.com/countertop-corrosives-acid-safety-cabinet-cap-six-2-1-2-liter-bottles-1-door-poly-white-24004 |  |  |  | https://www.justrite.com/media/catalog/product/_/2/_2_4_24004_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/2/_2_4_24004_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 12 Gal (Self-Close) | 891222 | $2,044.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Countertop and Compac Cabinets let you customize your work areas and make the most of available space. Whether adding to existing storage space or making chemicals readily available at workstation locations, these smaller-sized cabinets are the solution when space is at a premium.The ChemCor® thermoplastic coating on all interior surfaces provides superior protection against corrosion from drips and off-gassing from harsh chemicals. Non-porous surface resists staining and cleans up easily.ChemCor® Lined Compac Acid Safety Cabinets have the same quality features found in Justrite's flammable cabinets — including double-wall, 18-gauge (1mm) welded steel construction, 2" (51mm) liquid-tight containment sump, dual vents, adjustable shelves, three-point stainless steel bullet self-latching doors with U•Loc™ padlockable handle and Haz-Alert™ reflective warning labels. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance. Interior ChemCor® liner ensures complete corrosion resistance and protection. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.Self Close models meet OSHA, NFPA 30, NFPA 1 and the International Fire Code. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 101.0 | 18 | 35 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29299/name/12_Gallon%252C_1_Shelf%252C_1_Door%252C_Self_Close%252C_CorrosivesAcids_Safety_Cabinet%252C_ChemCor%25C2%25AE_Compac%252C_Blue_-_8912222.pdf/ |  |  | https://www.justrite.com/chemcor-compac-corrosives-acids-safety-cabinet-cap-12-gallons-1-shelf-1-s-c-door-blue-8912222 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_8912222_similar-func-equiv_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_8912222_similar-func-equiv_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 90 gal (Self-Close), Blue | 899022 | $3,100.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Justrite blue steel safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Self Close models meet OSHA, NFPA 30, NFPA 1 and the International Fire Code. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 387.0 | 34 | 65 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29271/name/90_Gallon%252C_2_Shelves%252C__2_Doors%252C_Self_Close%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Blue_-_899022.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-corrosives-acid-steel-safety-cabinet-cap-90-gallons-2-shelves-2-s-c-doors-blue-899022 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_899022_sfe-lg_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_899022_sfe-lg_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Blue (45 gal, Manual) | 894502 | $1,830.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Justrite blue steel safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Manual close models meet OSHA and NFPA 30. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 268.5 | 18 | 65 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29264/name/45_Gallon%252C_2_Shelves%252C_2_Doors%252C_Manual_Close%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Blue_-_894502.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-corrosives-acid-steel-safety-cabinet-cap-45-gal-2-shelves-2-m-c-doors-blue-894502 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_894502_2_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_894502_2_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Blue Steel 22 Gal (Manual) | 892302 | $1,452.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Safely store corrosives under a workbench or lab counter for quick and easy access with these Justrite blue steel Undercounter safety cabinets for corrosives, which have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Manual close models meet OSHA and NFPA 30. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 154.0 | 22 | 35 | 35 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29283/name/22_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_CorrosivesAcid_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Under_Counter%252C_Blue_-_892302.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-undercounter-corrosives-acid-stl-safety-cabinet-cap-22-gal-1-shlf-2-m-c-drs-blue-892302 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_892302_use-lg.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_892302_use-lg.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Blue Steel 22 Gal (Self-Close) | 892322 | $1,750.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Safely store corrosives under a workbench or lab counter for quick and easy access with these Justrite blue steel Undercounter safety cabinets for corrosives, which have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Self Close models meet OSHA, NFPA 30, NFPA 1 and the International Fire Code. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 158.0 | 22 | 35 | 35 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29284/name/22_Gallon%252C_1_Shelf%252C_2_Doors%252C_Self_Close%252C_CorrosivesAcid_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Under_Counter%252C_Blue_-_892322.pdf/ |  |  | https://www.justrite.com/892322-chemcor-undercounter-corrosives-acids-safety-cabinet-blue-892322 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_892322_sfe_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_892322_sfe_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Blue Steel 30 Gal (Manual) | 893002 | $1,433.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Justrite blue steel safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. This popular 30-gallon (114-L) cabinet also includes an extra polyethylene work tray that can be secured to cabinet top for a handy work surface.An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Manual close models meet OSHA and NFPA 30. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 196.5 | 18 | 44 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29260/name/30_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Blue_-_893002.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-corrosives-acid-stl-safety-cabinet-dims-44-h-cap-30-gal-1-shlf-2-m-c-drs-blue-893002 |  |  |  | https://www.justrite.com/media/catalog/product/8/9/893002_1-1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/8/9/893002_1-1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Blue Steel 30 gal (Slide/Self-Close) | 893082 | $2,376.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Justrite blue steel safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Self Close models meet OSHA, NFPA 30, NFPA 1 and the International Fire Code. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 224.0 | 18 | 44 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29262/name/30_Gallon%252C_1_Shelf%252C_1_Bi-Fold_Self-Close_Door%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Blue_-_893082.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-corrosives-acid-stl-safety-cabinet-cap-30-gal-1-shelf-1-bi-fold-s-c-door-blue-893082 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_894502-lg_14_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_894502-lg_14_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Blue Steel 4 Gal (Manual) | 890402 | $826.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Countertop and Compac Cabinets let you customize your work areas and make the most of available space. Whether adding to existing storage space or making chemicals readily available at workstation locations, these smaller-sized cabinets are the solution when space is at a premium.Justrite blue steel safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Manual close models meet OSHA and NFPA 30. Cabinets come with a ten-year warranty. Read More | [19 empty] | 61.0 | 17 | 22 | 17 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29277/name/4_Gallon%252C_1_Shelf%252C_1_Door%252C_Manual_Close%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Countertop%252C_Blue_-_890402.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-countertop-corrosives-acid-steel-safety-cabinet-cap-4-gal-1-shlf-1-m-c-dr-blue-890402 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_890402-lg_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_890402-lg_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Blue Steel 4 Gal (Self-Close) | 890422 | $1,010.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Countertop and Compac Cabinets let you customize your work areas and make the most of available space. Whether adding to existing storage space or making chemicals readily available at workstation locations, these smaller-sized cabinets are the solution when space is at a premium.Justrite blue steel safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Self Close models meet OSHA, NFPA 30, NFPA 1 and the International Fire Code. Cabinets come with a ten-year warranty. Read More | [19 empty] | 65.0 | 17 | 22 | 17 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29278/name/4_Gallon%252C_1_Shelf%252C_1_Door%252C_Self_Close%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Countertop%252C_Blue_-_890422.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-countertop-corrosives-acid-steel-safety-cabinet-cap-4-gal-1-shlf-1-s-c-dr-blue-890422 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_890422_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_890422_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Blue Steel 45 gal (Self-Close) | 894522 | $2,100.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Justrite blue steel safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents. Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Self Close models meet OSHA, NFPA 30, NFPA 1 and the International Fire Code. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 272.5 | 18 | 65 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29265/name/45_Gallon%252C_2_Shelves%252C_2_Doors%252C_Self_Close%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Blue_-_894522.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-corrosives-acid-steel-safety-cabinet-cap-45-gal-2-shelves-2-s-c-doors-blue-894522 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_894522_similar_20func_20equiv_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_894522_similar_20func_20equiv_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Piggyback 17 Gal (Self-Close) | 891722 | $1,570.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices.Designed to mount on top of our standard 30- and 45-gallon (114- and 170-L) safety cabinets, these Piggyback cabinets create a safe storage space without taking up any valuable floor space. Cabinets can also be used independently as a wall-mount unit or on a countertop.Justrite blue steel Piggyback safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Self Close models meet OSHA, NFPA 30, NFPA 1 and the International Fire Code. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 131.5 | 18 | 24 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29288/name/17_Gallon%252C_1_Shelf%252C_2_Doors%252C_Self_Close%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Piggyback%252C_Blue_-_891722.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-piggyback-corrosives-acid-steel-safety-cabinet-cap-17-gal-1-shlf-2-s-c-doors-blu-891722 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891722-lg_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891722-lg_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Piggyback Blue Steel 17 Gal (Manual) | 891702 | $1,340.00 | Research environments using many different types of chemicals have special storage requirements. Whether using acids, bases or solvents, it is critical that incompatible liquids are segregated to avoid disastrous explosions or fire. The SDS (Safety Data Sheet) should be reviewed to determine chemical characteristics and recommended storage practices.Designed to mount on top of our standard 30- and 45-gallon (114- and 170-L) safety cabinets, these Piggyback cabinets create a safe storage space without taking up any valuable floor space. Cabinets can also be used independently as a wall-mount unit or on a countertop.Justrite blue steel Piggyback safety cabinets for corrosives have the same quality safety features as the cabinets for flammables. All models come with double-wall construction, dual vents, grounding wire connections, adjustable shelves, leakproof sills, three-point self-latching doors and leveling feet. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed.Additionally, they include polyethylene trays attached to galvanized steel shelves and a separate polyethylene liner for the bottom sump resist aggressive chemicals. Easily remove the liner for quick cleaning of drips and leaks. An all epoxy baked-on powder-coat finish, inside and out, provides increased chemical resistance.Manual close models meet OSHA and NFPA 30. Cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 127.5 | 18 | 24 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29287/name/17_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_CorrosivesAcid_Steel_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Piggyback%252C_Blue_-_891702.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-piggyback-corrosives-acid-steel-safety-cabinet-cap-17-gal-1-shlf-2-m-c-doors-blu-891702 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891702-lg_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891702-lg_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip 120 gallon-Self Latching 2 door | 899270 | $6,436.00 | For storage of larger quantities of flammable solvents, the heavy-duty Drum Safety Cabinet is available in many styles and sizes. Built to the same high standards as the smaller Classic Safety Cabinets, a reinforced liquid-tight sump contains spills. An array of drum accessories and equipment is available to ensure safety and convenience when storing and handling drums of hazardous materials.Double-duty yellow cabinets safely store flammable liquids. They feature two separate compartments that fit one vertical 55- or 30-gallon (200- or 110-L) drum and up to twelve 5-gallon (19-L) safety cans. Cabinets feature self-latching, three-point locking door for maximum protection under fire conditions. Latching system engages automatically when the door is pushed closed.Cabinets include drum rollers in the sump and one half-depth shelf in the drum compartment plus two full-size adjustable shelves in the can compartment. Sturdy 18-gauge (1-mm) double-wall steel construction offers excellent durability. Self-close model features a fusible link that holds doors wide open during use and melts at 165°F (74ºC), allowing the doors to close automatically in the event of a fire.Interior Dimensions:Drum Side - 30.625"W x 60.5"H x 30.34"DCan Side - 23.25"W x 60.5"H x 30.34"D Read More | [19 empty] | 614.0 | 34 | 65 | 59 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29208/name/115_Gallon%252C_1_Drum%252C_3_Shelves%252C_2_Doors%252C_Self_Close%252C_Double_Duty_Flammable_Cabinet_with_Drum_Rollers%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_899270.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-dbl-duty-safety-cabinet-w-drm-rlrs-partition-store-drum-can-3-shlvs-2-s-c-drs-yel-899270 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_899270-sfe-md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_899270-sfe-md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX (5.2 Gal., Red) | 893001 | $1,307.00 | #html-body [data-pb-style=U650NTR],#html-body [data-pb-style=YEA3SFS]{background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll}#html-body [data-pb-style=U650NTR]{justify-content:flex-start;display:flex;flex-direction:column}#html-body [data-pb-style=YEA3SFS]{align-self:stretch}#html-body [data-pb-style=M5GMX66]{display:flex;width:100%}#html-body [data-pb-style=HIQLFOL]{justify-content:flex-start;display:flex;flex-direction:column;background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll;width:100%;align-self:stretch}One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Manual Close Doors (model 893001) is designed to safely store flammable liquids. Its exterior dimensions are 44 in H x 43 in W x 18 in D (1118 mm H x 1092 mm W x 457 mm D). These red cabinets can store 30 gal (114 L) of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight. Model 893001 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29937. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity. Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 893001 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements.Safety Considerations for Standard Safety CabinetsThe requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more. For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 192.7 | 18 | 44 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/28982/name/30_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Red_-_893001.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-flammable-safety-cabinet-dims-44-h-cap-30-gal-1-shelf-2-m-c-doors-red-893001 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_893000_1_new_md_1_4.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_893000_1_new_md_1_4.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 12 Gal (Manual) | 891200 | $856.00 | The Justrite Sure-Grip® EX Compac Safety Cabinet with a Manual Close Door (model 891200) is for the safe and code-compliant storage of solvents and flammable liquids, such as aerosols, cleaners, and degreasers, on countertops and workbenches. Designed to increase accessibility without compromising safety, this yellow cabinet is ideal for maintenance garages with limited floor space. It is available in a portable size with exterior dimensions of 35 in H x 23.25 in W x 18 in D (889 mm H x 591 mm W x 457 mm D). Model 891200 can store 12 gal (45 L) of flammable liquids. It easily sits on countertops and fits under most workbenches. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, manual-close, self-latching door (fail safe, 3-point stainless steel) opens to a full 180 degrees and closes securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the door closes smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 891200 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29936.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891200 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Countertop Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 103.0 | 18 | 35 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29079/name/12_Gallon%252C_1_Shelf%252C_1_Door%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Compac%252C_Yellow_-_891200.pdf/ |  |  | https://www.justrite.com/891200-sure-grip-ex-compac-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891200-1-md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891200-1-md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 12 Gal (Self-Close) | 891220 | $1,020.00 | The Justrite Sure-Grip® EX Compac Safety Cabinet with a Self-Close Door (model 891220) is for the safe and code-compliant storage of solvents and flammable liquids, such as aerosols, cleaners, and degreasers, on countertops and workbenches. Designed to increase accessibility without compromising safety, this yellow cabinet is ideal for maintenance garages with limited floor space. It is available in a portable size with exterior dimensions of 35 in H x 23.25 in W x 18 in D (889 mm H x 591 mm W x 457 mm D). Model 891220 can store 12 gal (45 L) of flammable liquids. It easily sits on countertops and fits under most workbenches. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, self-closing door opens to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. Continuous piano hinges ensure the door closes smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 891220 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29936.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891220 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1 and IFC requirements. Safety Considerations for Countertop Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 100.0 | 18 | 35 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29083/name/12_Gallon%252C_1_Shelf%252C_1_Door%252C_Self_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Compac%252C_Yellow_-_891220.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-compac-flammable-safety-cabinet-cap-12-gallons-1-shelf-1-s-c-door-yellow-891220 |  |  |  | https://www.justrite.com/media/catalog/product/_/s/_s_u_super-grip-ex-compac-891220_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/s/_s_u_super-grip-ex-compac-891220_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 15 Gal (Manual) | 891500 | $965.00 | The Justrite Sure-Grip® EX Compac Safety Cabinet with a Manual Close Door (model 891500) is for the safe and code-compliant storage of solvents and flammable liquids, such as aerosols, cleaners, and degreasers, on countertops and workbenches. Designed to increase accessibility without compromising safety, this yellow cabinet is ideal for maintenance garages with limited floor space. It is available in a portable size with exterior dimensions of 44 in H x 23.25 in W x 18 in D (1118 mm H x 591 mm W x 457 mm D). Model 891500 can store 15 gal (57 L) of flammable liquids and easily sits on countertops. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, manual-close, self-latching door (fail safe, 3-point stainless steel) opens to a full 180 degrees and closes securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the door closes smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 891500 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29936.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891500 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Countertop Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 116.0 | 18 | 44 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29087/name/15_Gallon%252C_1_Shelf%252C_1_Door%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Compac%252C_Yellow_-_891500.pdf/ |  |  | https://www.justrite.com/891500-sure-grip-ex-compac-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891500_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891500_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 20 Gal (Manual) | 893400 | $1,273.00 | Gain storage space near work areas without taking up extra floor space. Eye-level positioning allows for easy access to solvents. Model 893400 features three adjustable shelves to accommodate multiple containers. Both sizes come complete with mounting hardware. These cabinets have all the great features and benefits of our classic cabinets as outlined below. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Protect workers, reduce fire risks, and improve productivity by storing flammable fuels and chemicals in code-compliant safety cabinets. Designed to meet OSHA and NFPA standards, Sure-Grip® EX cabinets are constructed of sturdy 18-gauge (1-mm) thick double-wall, welded steel with 1-1/2” (38-mm) of insulating air space for fire resistance. High-performance, manual-close, self-latching doors open to a full 180 degrees and close easily and securely for maximum protection under fire conditions when pushed shut. Fail-safe closing mechanism ensures three-point stainless steel bullet latching system works every time. For added security, U-Loc™ handle comes with a cylinder lock and key set or accepts an optional padlock. Haz-Alert™ reflective warning labels are highly visible under fire conditions or during power outages when illuminated with a flashlight.Patented SpillSlope® galvanized steel shelves direct spills to back and bottom of a leakproof 2” (51-mm) bottom sump. They adjust on 3” (76-mm) centers for versatile storage. Cabinets also feature dual vents with flame arresters, four adjustable self-leveling feet, grounding connector, and trilingual warning label. Durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance.All cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 188.5 | 12 | 44 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29124/name/20_Gallon%252C_3_Shelves%252C_2_Doors%252C_Manual_Close%252C_Wall_Mount_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_893400.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-wall-mount-flammable-safety-cabinet-cap-20-gallons-5-shelves-2-m-c-doors-yellow-893400 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_893400-md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_893400-md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 20 Gal Red (Self-Close) | 891531 | $1,400.00 | Easily separate and identify Class III combustibles using these red safety cabinets. Adjustable shelves accommodate many smaller paint tins, aerosols and inks found in printing, painting, furniture and motor vehicle operations. Sure-Grip® EX combustibles storage cabinets have all the same quality features as Justrite’s flammable cabinets as outlined below.Designed to meet OSHA and NFPA standards, Sure-Grip® EX cabinets are constructed of sturdy 18-gauge (1-mm) thick double-wall, welded steel with 1-1/2 in. (38-mm) of insulating air space for fire resistance. High-performance, self-latching doors close easily and securely for maximum protection under fire conditions. Fail-safe closing mechanism ensures three-point stainless steel bullet latching system works every time.For added security, U-Loc™ handle comes with a cylinder lock and key set or accepts an optional padlock. Haz-Alert™ reflective warning labels are highly visible under fire conditions or during power outages when illuminated with a flashlight.Patented SpillSlope® galvanized steel shelves direct spills to back and bottom of a leakproof sump. They adjust on 3 in. (76-mm) centers for versatile storage. Cabinets also feature a leakproof 2 in. (51-mm) bottom sump, dual vents with flame arresters, four adjustable self-leveling feet, grounding connector and trilingual warning label. Durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance.Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.All door styles meet OSHA and NFPA 30; self-close door styles also meet NFPA 1 and the International Fire Code.All cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 139.0 | 18 | 44 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29221/name/20_Gallon%252C_2_Shelves%252C_1_Door%252C_Self_Close%252C_Paint_Safety_Cabinet%252C__Sure-Grip%25C2%25AE_EX%252C_Red_-_891531.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-combustibles-safety-cabinet-for-paint-and-ink-cap-20-gal-2-shlves-1-s-c-door-red-891531 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891531_similar_20func_20equiv_md_1_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891531_similar_20func_20equiv_md_1_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 22 Gal (Manual) | 892300 | $1,136.00 | Conveniently locate flammable liquids near work areas, with the Justrite Sure-Grip® Undercounter Safety Cabinet with Manual Close Doors (model 892300). It can be safely placed beneath countertops and work benches and features a 3.5 in H x 2.2 in D (89 mm H x 64 mm D) recessed toe kick. This allows users to stand closer to the work surface, which means less leaning and better ergonomics. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Its exterior dimensions are 35 in x 35 in x 22 in (889 mm x 889 mm x 559 mm). These yellow cabinets can store 22 gal (83 l) of flammable liquids.This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 892300 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For a replacement shelf, use model 29939. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet that adjust to 1.125 in (29 mm) for alignment on uneven floors, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-l), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 892300 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Undercounter Safety Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 163.0 | 22 | 35 | 35 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29095/name/22_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C__Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Under_Counter%252C_Yellow_-_892300.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-undercounter-flammable-safety-cabinet-cap-22-gallons-1-shelf-2-m-c-doors-yellow-892300 |  |  |  | https://www.justrite.com/media/catalog/product/8/9/892300_2-1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/8/9/892300_2-1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 22 Gal (Self-Close) | 892320 | $1,405.00 | Conveniently locate flammable liquids near work areas, with the Justrite Sure-Grip® Undercounter Safety Cabinet with Self Close Doors (model 892320). It can be safely placed beneath countertops and work benches and features a 3.5 in H x 2.2 in D (89 mm H x 64 mm D) recessed toe kick. This allows users to stand closer to the work surface, which means less leaning and better ergonomics. Its exterior dimensions are 35 in x 35 in x 22 in (889 mm x 889 mm x 559 mm). These yellow cabinets can store 22 gal (83 l) of flammable liquids. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, self-closing doors open to a full 180 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tightly sequence. Continuous piano hinges ensure the doors close smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 892320 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For a replacement shelf, use model 29939. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet that adjust to 1.125 in (29 mm) for alignment on uneven floors, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-l), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 892320 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements. Safety Considerations for Undercounter Safety Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 152.0 | 22 | 35 | 35 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29099/name/22_Gallon%252C_1_Shelf%252C_2_Doors%252C_Self_Close%252C__Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Under_Counter%252C_Yellow_-_892320.pdf/ |  |  | https://www.justrite.com/892320-sure-grip-ex-undercounter-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_892320_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_892320_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 22 Gal, 1 Self-Close Doors | 892225 | $1,520.00 | Gain valuable storage space even when floor space is tight. The Justrite Sure-Grip® EX Slimline Safety Cabinet with 1 Self Close Door (model 892225) is designed for the safe storage and segregation of flammable liquids in lab, construction, or engineering settings where floor space is limited. It fits flush alongside a 45 gal (170 L) cabinet, which makes it easy to help expand storage alongside an existing cabinet system. The cabinet’s smaller footprint allows placement in labs, R&D rooms, garages, machine shops, and more. Its exterior dimensions are 65 in H x 23.25 in W x 18 in D (1651 mm H x 591 mm W x 457 mm D). These white cabinets can store 22 gal (83 L) of flammable liquids. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, self-closing door can open to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. Continuous piano hinges ensure the door closes smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 892225 features three patented, galvanized steel SpillSlope® shelves designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29936. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 892225 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements. Safety Considerations for Justrite Slimline Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 191.0 | 18 | 65 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29061/name/22_Gallon%252C_3_Shelves%252C_1_Door%252C_Self_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Slimline%252C_White_-_892225.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-slimline-flammable-safety-cabinet-cap-22-gallons-3-shelves-1-s-c-door-white-892225 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_892225_4.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_892225_4.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 24 Aerosol Cans, 2 Drawers (Manual) | 890500 | $1,490.00 | Justrite's Aerosol Can Benchtop Safety Cabinet fits virtually anywhere for hassle-free organization of up to twenty-four aerosol cans. Cabinet features 24-gauge galvanized steel roll-out shelves with "easy-glide" plastic rollers and stop guard protection. Two fixed position shelves, with twelve 2¾ in. (70mm) diameter openings, accept cans up to 10 in. (254mm) tall and are removable for easy cleaning. These cabinets have all the great features and benefits of our classic cabinets as outlined below. Protect workers, reduce fire risks, and improve productivity by storing flammable fuels and chemicals in code-compliant safety cabinets. Designed to meet OSHA and NFPA standards, Sure-Grip® EX cabinets are constructed of sturdy 18-gauge (1-mm) thick double-wall, welded steel with 1-1/2 in. (38-mm) of insulating air space for fire resistance. High-performance, self-latching doors close easily and securely for maximum protection under fire conditions. Fail-safe closing mechanism ensures three-point stainless steel bullet latching system works every time. Looking for a custom safety cabinet? Request a quote here for your storage requirements. For added security, U-Loc™ handle comes with a cylinder lock and key set or accepts an optional padlock. Haz-Alert™ reflective warning labels are highly visible under fire conditions or during power outages when illuminated with a flashlight. Cabinets feature a leakproof 2 in. (51-mm) bottom sump, four adjustable self-leveling feet, grounding connector and trilingual warning label. Durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance.Cabinets come with a ten-year warranty. Read More | [19 empty] | 91.0 | 18 | 27 | 21 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29125/name/1_Door%252C_Manual_Close%252C_2_Drawers%252C_24_Can_Benchtop_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_890500.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-benchtop-flammable-safety-cabinet-cap-24-aerosol-cans-2-drawers-1-m-c-door-yel-890500 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_890500-md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_890500-md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 30 Gal (Manual) | 893000 | $1,144.00 | #html-body [data-pb-style=QLL5VDI],#html-body [data-pb-style=WFQSCLG]{background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll}#html-body [data-pb-style=QLL5VDI]{justify-content:flex-start;display:flex;flex-direction:column}#html-body [data-pb-style=WFQSCLG]{align-self:stretch}#html-body [data-pb-style=LK9594W]{display:flex;width:100%}#html-body [data-pb-style=XNWFRCG]{justify-content:flex-start;display:flex;flex-direction:column;background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll;width:100%;align-self:stretch}One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Manual Close Doors (model 893000) is designed to safely store flammable liquids. Its exterior dimensions are 44 in H x 43 in W x 18 in D (1118 mm H x 1092 mm W x 457 mm D). These yellow cabinets can store 30 gal (114 l) of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight. Model 893000 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use models 29943 or 29937. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity. Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-l), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 893000 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements.Safety Considerations for Standard Safety CabinetsThe requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants and more. For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 192.7 | 18 | 44 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/28981/name/30_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_893000.pdf/ |  |  | https://www.justrite.com/893000-sure-grip-ex-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/8/9/893000_5.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/8/9/893000_5.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 30 gal (Manual) | 893300 | $1,360.00 | #html-body [data-pb-style=RNE3AML],#html-body [data-pb-style=TAGEAJ9]{background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll}#html-body [data-pb-style=RNE3AML]{justify-content:flex-start;display:flex;flex-direction:column}#html-body [data-pb-style=TAGEAJ9]{align-self:stretch}#html-body [data-pb-style=C10B6DX]{display:flex;width:100%}#html-body [data-pb-style=J7APYRB]{justify-content:flex-start;display:flex;flex-direction:column;background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll;width:100%;align-self:stretch}One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Manual Close Doors (model 893300) is designed to safely store flammable liquids. Its exterior dimensions are 35 in H x 36 in W x 24 in D (889 mm H x 914 mm W x 610 mm D). These yellow cabinets can store 30 gal (114 L) of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight. Model 893300 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29942. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity. Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 893300 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements.Safety Considerations for Standard Safety CabinetsThe requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more. For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 194.0 | 24 | 35 | 36 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29025/name/30_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_35%2522H_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_893300.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-flammable-safety-cabinet-dims-35-h-cap-30-gal-1-shelf-2-m-c-doors-yellow-893300 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_893300_similar_20func_20equiv_md_4_5.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_893300_similar_20func_20equiv_md_4_5.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 30 Gal (Self-Close) | 893020 | $1,304.00 | #html-body [data-pb-style=CU5C0S1],#html-body [data-pb-style=ET0EBIO]{background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll}#html-body [data-pb-style=CU5C0S1]{justify-content:flex-start;display:flex;flex-direction:column}#html-body [data-pb-style=ET0EBIO]{align-self:stretch}#html-body [data-pb-style=TWWDJUS]{display:flex;width:100%}#html-body [data-pb-style=QSS3WTE]{justify-content:flex-start;display:flex;flex-direction:column;background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll;width:100%;align-self:stretch}One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Self-Close Doors (model 893020) is designed to safely store flammable liquids. Its exterior dimensions are 44 in H x 43 in W x 18 in D (1118 mm H x 1092 mm W x 457 mm D). These yellow cabinets can store 30 gal (114 l) of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, self-closing doors open to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tightly sequence. Continuous piano hinges ensure the doors close smoothly every time. When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight. Model 893020 features one patented, galvanized steel SpillSlope® shelf designed to directs spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use models 29943 or 29937. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity. Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435- l), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 893020 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements.Safety Considerations for Standard Safety CabinetsThe requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants and more. For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 194.5 | 18 | 44 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29020/name/30_Gallon%252C_1_Shelf%252C_2_Doors%252C_Self_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_893020.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-flammable-safety-cabinet-cap-30-gallons-1-shelf-2-self-close-doors-yellow-893020 |  |  |  | https://www.justrite.com/media/catalog/product/8/9/893020_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/8/9/893020_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 4 Gal (Manual) | 890400 | $681.00 | The Justrite Sure-Grip® EX Countertop Safety Cabinet with a Manual Close Door (model 890400) is for the safe and code-compliant storage of solvents and flammable liquids such as aerosols, cleaners, and degreasers, on countertops and workbenches. Designed to increase accessibility without compromising safety, this yellow cabinet is ideal for maintenance garages with limited floor space. It is available in a portable size with exterior dimensions of 22 in H x 17 in W x 17 in D (559 mm H x 432 mm W x 432 mm D). Model 890400 can store 4 gal (15l) of flammable liquids. It easily sits on countertops and is also wall-mountable with an optional wall hanger. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, manual-close, self-latching door (fail safe, 3-point stainless steel) opens to a full 180 degrees and closes securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the door closes smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 890400 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29935.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-l), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 890400 complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Countertop Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 54.2 | 17 | 22 | 17 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29071/name/4_Gallon%252C_1_Shelf%252C_1_Door%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Countertop%252C_Yellow_-_890400.pdf/ |  |  | https://www.justrite.com/890400-sure-grip-ex-countertop-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_890400-md_4.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_890400-md_4.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 4 Gal (Manual) | 890401 | $745.00 | The Justrite Sure-Grip® EX Countertop Safety Cabinet with a Manual Close Door (model 890401) is for the safe and code-compliant storage of solvents and flammable liquids, such as aerosols, cleaners, and degreasers, on countertops and workbenches. Designed to increase accessibility without compromising safety, this red cabinet is ideal for maintenance garages with limited floor space. It is available in a portable size with exterior dimensions of 22 in H x 17 in W x 17 in D (559 mm H x 432 mm W x 432 mm D). Model 890401 can store 4 gal (15 L) of flammable liquids. It easily sits on countertops and is also wall-mountable with an optional wall hanger.This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. The high-performance, manual-close, self-latching door (fail safe, 3-point stainless steel) opens to a full 180 degrees and closes securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the door closes smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 890401 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29935.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 890401 complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Countertop Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 56.0 | 17 | 22 | 17 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29072/name/4_Gallon%252C_1_Shelf%252C_1_Door%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Countertop%252C_Red_-_890401.pdf/ |  |  | https://www.justrite.com/890401-sure-grip-ex-countertop-flammable-safety-cabinet-red |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_890401_4.png?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_890401_4.png?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 4 Gal (Self-Close) | 890420 | $776.00 | The Justrite Sure-Grip® EX Countertop Safety Cabinet with a Self-Close Door (model 890420) is for the safe and code-compliant storage of solvents and flammable liquids such as aerosols, cleaners, and degreasers, on countertops and workbenches. Designed to increase accessibility without compromising safety, this yellow cabinet is ideal for maintenance garages with limited floor space. It is available in a portable size with exterior dimensions of 22 in H x 17 in W x 17 in D (559 mm H x 432 mm W x 432 mm D). Model 890420 can store 4 gal (15l) of flammable liquids. It easily sits on countertops. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, self-closing door opens to a full 180 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. Continuous piano hinges ensure the doors close smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 890420 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29935.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-l), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 890420 complies with OSHA, NFPA 30, NFPA 1 and IFC requirements. Safety Considerations for Countertop Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 59.0 | 17 | 22 | 17 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29075/name/4_Gallon%252C_1_Shelf%252C_1_Door%252C_Self_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Countertop%252C_Yellow_-_890420.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-countertop-flammable-safety-cabinet-cap-4-gallons-1-shelf-1-s-c-door-yellow-890420 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_890420_use_md.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_890420_use_md.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 40 Gal (Red/Manual) | 893011 | $1,383.00 | Easily separate and identify Class III combustibles using these red safety cabinets. Adjustable shelves accommodate many smaller paint tins, aerosols and inks found in printing, painting, furniture and motor vehicle operations. Sure-Grip® EX combustibles storage cabinets have all the same quality features as Justrite’s flammable cabinets as outlined below.Designed to meet OSHA and NFPA standards, Sure-Grip® EX cabinets are constructed of sturdy 18-gauge (1-mm) thick double-wall, welded steel with 1-1/2 in. (38-mm) of insulating air space for fire resistance. High-performance, self-latching doors close easily and securely for maximum protection under fire conditions. Fail-safe closing mechanism ensures three-point stainless steel bullet latching system works every time.For added security, U-Loc™ handle comes with a cylinder lock and key set or accepts an optional padlock. Haz-Alert™ reflective warning labels are highly visible under fire conditions or during power outages when illuminated with a flashlight.Patented SpillSlope® galvanized steel shelves direct spills to back and bottom of a leakproof sump. They adjust on 3 in. (76-mm) centers for versatile storage. Cabinets also feature a leakproof 2 in. (51-mm) bottom sump, dual vents with flame arresters, four adjustable self-leveling feet, grounding connector and trilingual warning label. Durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance.Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.All door styles meet OSHA and NFPA 30; self-close door styles also meet NFPA 1 and the International Fire Code.All cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 240.0 | 18 | 44 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29223/name/40_Gallon%252C_3_Shelves%252C_2_Doors%252C_Manual_Close%252C_Paint_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Red_-_893011.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-combustibles-safety-cabinet-for-paint-and-ink-cap-40-gal-3-shelves-2-m-c-dr-red-893011 |  |  |  | https://www.justrite.com/media/catalog/product/8/9/893011_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/8/9/893011_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 40 Gal (Self-Close) | 893030 | $1,552.00 | Easily separate and identify Class III combustibles using these red safety cabinets. Adjustable shelves accommodate many smaller paint tins, aerosols and inks found in printing, painting, furniture and motor vehicle operations. Sure-Grip® EX combustibles storage cabinets have all the same quality features as Justrite’s flammable cabinets as outlined below.Designed to meet OSHA and NFPA standards, Sure-Grip® EX cabinets are constructed of sturdy 18-gauge (1-mm) thick double-wall, welded steel with 1-1/2 in. (38-mm) of insulating air space for fire resistance. High-performance, self-latching doors close easily and securely for maximum protection under fire conditions. Fail-safe closing mechanism ensures three-point stainless steel bullet latching system works every time.For added security, U-Loc™ handle comes with a cylinder lock and key set or accepts an optional padlock. Haz-Alert™ reflective warning labels are highly visible under fire conditions or during power outages when illuminated with a flashlight.Patented SpillSlope® galvanized steel shelves direct spills to back and bottom of a leakproof sump. They adjust on 3 in. (76-mm) centers for versatile storage. Cabinets also feature a leakproof 2 in. (51-mm) bottom sump, dual vents with flame arresters, four adjustable self-leveling feet, grounding connector and trilingual warning label. Durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance.Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.All door styles meet OSHA and NFPA 30; self-close door styles also meet NFPA 1 and the International Fire Code.All cabinets are FM approved and come with a ten-year warranty. Read More | [19 empty] | 200.0 | 18 | 44 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29224/name/40_Gallon%252C_3_Shelves%252C_2_Doors%252C_Self_Close%252C_Paint_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_893030.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-combustibles-safety-cabinet-for-paint-and-ink-cap-40-gal-3-shlves-2-s-c-door-yel-893030 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_893030_sfe_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_893030_sfe_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 45 gal (Manual) | 894500 | $1,465.00 | #html-body [data-pb-style=OY8S6HV],#html-body [data-pb-style=SD2QWLQ]{background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll}#html-body [data-pb-style=OY8S6HV]{justify-content:flex-start;display:flex;flex-direction:column}#html-body [data-pb-style=SD2QWLQ]{align-self:stretch}#html-body [data-pb-style=ERTJGLG]{display:flex;width:100%}#html-body [data-pb-style=KUEOPCE]{justify-content:flex-start;display:flex;flex-direction:column;background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll;width:100%;align-self:stretch}One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Manual Close Doors (model 894500) is designed to safely store flammable liquids. Its exterior dimensions are 65 in H x 43 in W x 18 in D (1651 mm H x 1092 mm W x 457 mm D). These yellow cabinets can store 45 gal (170 L) of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight. Model 894500 features two patented, galvanized steel SpillSlope® shelves designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29937. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity. Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 894500 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements.Safety Considerations for Standard Safety CabinetsThe requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more. For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Want to store flammable liquids outdoors? Check out the 45-gallon outdoor flammable storage cabinet. Read More | [19 empty] | 264.7 | 18 | 65 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29029/name/45_Gallon%252C_2_Shelves%252C_2_Doors%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_894500.pdf/ |  |  | https://www.justrite.com/894500-45-gallon-sure-grip-ex-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_894500_8_md_4_5.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_894500_8_md_4_5.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 45 gal (Self-Close) | 894520 | $1,758.00 | #html-body [data-pb-style=GQNJJVF],#html-body [data-pb-style=SJ7SH6A]{background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll}#html-body [data-pb-style=GQNJJVF]{justify-content:flex-start;display:flex;flex-direction:column}#html-body [data-pb-style=SJ7SH6A]{align-self:stretch}#html-body [data-pb-style=VK8CLFJ]{display:flex;width:100%}#html-body [data-pb-style=V6Y0YAH]{justify-content:flex-start;display:flex;flex-direction:column;background-position:left top;background-size:cover;background-repeat:no-repeat;background-attachment:scroll;width:100%;align-self:stretch}One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Self-Close Doors (model 894520) is designed to safely store flammable liquids. Its exterior dimensions are 65 in H x 43 in W x 18 in D (1651 mm H x 1092 mm W x 457 mm D). These yellow cabinets can store 45 gal (170 L) of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, self-closing doors open to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tight sequence. Continuous piano hinges ensure the doors close smoothly every time. When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight. Model 894520 features two patented, galvanized steel SpillSlope® shelves designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29937. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity. Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 894520 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements.Safety Considerations for Standard Safety CabinetsThe requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more. For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Want to store flammable liquids outdoors? Check out the 45-gallon outdoor flammable storage cabinet. Read More | [19 empty] | 266.5 | 18 | 65 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29033/name/45_Gallon%252C_2_Shelves%252C_2_Doors%252C_Self_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_894520.pdf/ |  |  | https://www.justrite.com/894520-45-gallon-sure-grip-ex-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/8/9/894520_4.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/8/9/894520_4.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 55 gal (Manual) | 896200 | $1,937.00 | For storage of larger quantities of flammable solvents, the heavy-duty Drum Safety Cabinet is available in many styles and sizes. Built to the same high standards as the smaller Classic Safety Cabinets, a reinforced liquid-tight sump contains spills. An array of drum accessories and equipment is available to ensure safety and convenience when storing and handling drums of hazardous materials.Use these vertical drum safety cabinets to safely store 30- to 55-gallon (110- to 200-L) drums. The removable, adjustable, half-depth shelf provides enough head space to accommodate a drum with a pump or funnel so the cabinet can be used as a pumping or waste accumulation station, and also allows you to store Class III materials or spare accessories. Available with or without roller assemblies to easily move heavy drums. FM approved. | [19 empty] | 346.0 | 34 | 65 | 34 |  |  |  | https://www.justrite.com/productpdf/download/file/id/28999/name/55_Gallon%252C_1_Drum_Vertical%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_Flammable_Cabinet_w_Drum_Support%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_896200.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-vertical-drum-safety-cabinet-and-drum-support-cap-55-gal-1-shlf-2-m-c-doors-yel-896200 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_896200_closed_5.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_896200_closed_5.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 60 gal (Manual) | 896000 | $1,768.00 | One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Manual Close Doors (model 896000) is designed to safely store flammable liquids. Its exterior dimensions are 65 in H x 34 in W x 34 in D (1651 mm H x 864 mm W x 864 mm D). These yellow cabinets can store 60 gal (227 L) of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 896000 features two patented, galvanized steel SpillSlope® shelves designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29944. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 896000 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Standard Safety Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 329.3 | 34 | 65 | 34 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29039/name/60_Gallon%252C_2_Shelves%252C_2_Doors%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_896000.pdf/ |  |  | https://www.justrite.com/896000-sure-grip-ex-flammable-safety-cabinet-cap-60-gallons-2-shelves-2-manual-close-doors-yellow |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_896000-sure-grip-flammable-safety-cabinet-60-gal-2-manual-doors-yellow_4_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_896000-sure-grip-flammable-safety-cabinet-60-gal-2-manual-doors-yellow_4_1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 60 gal (Self-Close) | 896020 | $2,000.00 | One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Self-Close Doors (model 896020) is designed to safely store flammable liquids. Its exterior dimensions are 65 in H x 34 in W x 34 in D (1651 mm H x 864 mm W x 864 mm D). These yellow cabinets can store 60 gal (227 L) of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, self-closing doors open to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tight sequence. Continuous piano hinges ensure the doors close smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read "FLAMMABLE – KEEP FIRE AWAY" are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 896020 features two patented, galvanized steel SpillSlope® shelves designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29944. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 896020 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements. Safety Considerations for Standard Safety Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 331.1 | 34 | 65 | 34 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29043/name/60_Gallon%252C_2_Shelves%252C_2_Doors%252C_Self_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_896020.pdf/ |  |  | https://www.justrite.com/896020-sure-grip-ex-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_896020-md_5.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_896020-md_5.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 90 gal (Manual) | 899000 | $2,105.00 | One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Manual Close Doors (model 899000) is designed to safely store flammable liquids. Its exterior dimensions are 65 in H x 43 in W x 34 in D (1651 mm H x 1092 mm W x 864 mm D). These yellow cabinets can store 90 gal (341 L)* of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 899000 features two patented, galvanized steel SpillSlope® shelves designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29945. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 899000 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Standard Safety Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage.*OSHA 29 CFR 1910.106(d)(3) states that not more than 60 gallons (227L) of Category 1, 2, or 3 flammable liquids, nor more than 120 gallons (454L) of Category 4 flammable liquids may be stored in a safety cabinet. document.addEventListener("DOMContentLoaded", function() { // Select the div with class "additional-info" const additionalInfoDiv = document.querySelector("div.additional-info"); if (additionalInfoDiv) { // Find the <img> element within this div const imgElement = additionalInfoDiv.querySelector("img[src='https://www.justrite.com/media/wysiwyg/made-in-usa.png']"); // If the <img> element is found, update the src if (imgElement) { imgElement.src = "https://www.justrite.com/media/wysiwyg/made-in-usa-new.png"; console.log("Image source updated successfully!"); } else { console.log("Image with the specified src not found."); } } else { console.log("Div with class 'additional-info' not found."); } }); Read More | [19 empty] | 379.2 | 34 | 65 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29049/name/90_Gallon%252C_2_Shelves%252C_2_Doors%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_899000.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-flammable-safety-cabinet-cap-90-gallons-2-shelves-2-manual-close-doors-yellow-899000 |  |  |  | https://www.justrite.com/media/catalog/product/8/9/899000_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/8/9/899000_1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX 90 gal (Self-Close) | 899020 | $2,500.00 | One of the leading causes of industrial fires is the improper storage and handling of flammable liquids. The Justrite Standard Sure-Grip® EX Flammable Safety Cabinet with Self-Close Doors (model 899020) is designed to safely store flammable liquids. Its exterior dimensions are 65 in H x 43 in W x 34 in D (1651 mm H x 1092 mm W x 864 mm D). These yellow cabinets can store 90 gal (341 L)* of flammable liquids. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. Looking for a custom safety cabinet? Request a quote here for your storage requirements. High-performance, self-closing doors open to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tight sequence. Continuous piano hinges ensure the doors close smoothly every time. When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet's self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Trilingual Haz-Alert™ reflective warning labels that read "FLAMMABLE – KEEP FIRE AWAY" are highly visible under fire conditions or during power outages when illuminated with a flashlight. Model 899020 features two patented, galvanized steel SpillSlope® shelves designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29945. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity. Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 899020 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements. Safety Considerations for Standard Safety Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more. For more information, please consult our workplace safety article about the requirements for flammable liquid storage. *OSHA 29 CFR 1910.106(d)(3) states that not more than 60 gallons (227L) of Category 1, 2, or 3 flammable liquids, nor more than 120 gallons (454L) of Category 4 flammable liquids may be stored in a safety cabinet. Read More | [19 empty] | 381.0 | 34 | 65 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29053/name/90_Gallon%252C_2_Shelves%252C_2_Doors%252C_Self_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_899020.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-flammable-safety-cabinet-cap-90-gallons-2-shelves-2-self-close-doors-yellow-899020 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_899020_md_4.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_899020_md_4.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Compac 15 Gal (Self-Close) | 891520 | $1,100.00 | The Justrite Sure-Grip® EX Compac Safety Cabinet with a Self-Close Door (model 891520) is for the safe and code-compliant storage of solvents and flammable liquids, such as aerosols, cleaners, and degreasers, on countertops and workbenches. Designed to increase accessibility without compromising safety, this yellow cabinet is ideal for maintenance garages with limited floor space. It is available in a portable size with exterior dimensions of 44 in H x 23.25 in W x 18 in D (1118 mm H x 591 mm W x 457 mm D). Model 891520 can store 15 gal (57 L) of flammable liquids and easily sits on countertops. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, self-closing door opens to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. Continuous piano hinges ensure the door closes smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 891520 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29936.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891520 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1 and IFC requirements. Safety Considerations for Countertop Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 118.5 | 18 | 44 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29091/name/15_Gallon%252C_1_Shelf%252C_1_Door%252C_Self_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Compac%252C_Yellow_-_891520.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-compac-flammable-safety-cabinet-cap-15-gallons-1-shelf-1-s-c-door-yellow-891520 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891520_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891520_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Compac Flammable Safety Cabinet 12 Ga | 891221 | $1,120.00 | The Justrite Sure-Grip® EX Compac Safety Cabinet with a Self-Close Door (model 891221) is for the safe and code-compliant storage of solvents and flammable liquids, such as aerosols, cleaners, and degreasers, on countertops and workbenches. Designed to increase accessibility without compromising safety, this red cabinet is ideal for maintenance garages with limited floor space. It is available in a portable size with exterior dimensions of 35 in H x 23.25 in W x 18 in D (889 mm H x 591 mm W x 457 mm D). Model 891221 can store 12 gal (45 L) of flammable liquids. It easily sits on countertops and fits under most workbenches. Looking for a custom safety cabinet? Request a quote here for your storage requirements. This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, self-closing door opens to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. Continuous piano hinges ensure the door closes smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 891221 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29936.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891221 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1 and IFC requirements. Safety Considerations for Countertop Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | N/A | 18 | 35 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29084/name/12_Gallon%252C_1_Shelf%252C_1_Door%252C_Self_Close%252C__Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Compac%252C_Red_-_891221.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-compac-flammable-safety-cabinet-cap-12-gallons-1-shelf-1-s-c-door-red-891221 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891221_3.png?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891221_3.png?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Flammable Safety Cabinet 90 Gallon | 899080 | $3,400.00 | Protect workers, reduce fire risks, and improve productivity by storing flammable fuels and chemicals in code-compliant safety cabinets. Designed to meet OSHA and NFPA standards, Sure-Grip® EX cabinets are constructed of sturdy 18-gauge (1-mm) thick double-wall, welded steel with 1-1/2" (38-mm) of insulating air space for fire resistance. High-performance, self-latching doors close easily and securely for maximum protection under fire conditions. Fail-safe closing mechanism ensures three-point stainless steel bullet latching system works every time.For added security, U-Loc™ handle comes with a cylinder lock and key set or accepts an optional padlock. Haz-Alert™ reflective warning labels are highly visible under fire conditions or during power outages when illuminated with a flashlight.Looking for a custom safety cabinet? Request a quote here for your storage requirements.Patented SpillSlope® galvanized steel shelves direct spills to back and bottom of a leakproof 2” (51-mm) bottom sump. They adjust on 3” (76-mm) centers for versatile storage. Cabinets also feature dual vents with flame arresters, four adjustable self-leveling feet, grounding connector, and trilingual warning label. Durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents.All door styles meet OSHA and NFPA 30; self-close door styles also meet NFPA 1 and the International Fire Code.Most cabinets are FM approved and all come with a ten-year warranty. Read More | [19 empty] | 433.0 | 34 | 65 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29057/name/90_Gallon%252C_2_Shelves%252C_1_Bi-Fold_Self_Close_Door%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_899080.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-flammable-safety-cabinet-cap-90-gallons-2-shelves-1-bi-fold-s-c-door-yellow-899080 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_899080_sfe_md_4.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_899080_sfe_md_4.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Mini 1 Gal (Manual) | 890200 | $622.00 | Mini Safety Cabinets offer the smallest footprint for easy mobility on service vehicles, carts and maintenance trikes. Includes 3 bungee cords to keep contents in place during movement. These cabinets have all the great features and benefits of our classic cabinets as outlined below. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Protect workers, reduce fire risks, and improve productivity by storing flammable fuels and chemicals in code-compliant safety cabinets. Designed to meet OSHA and NFPA standards, Sure-Grip® EX cabinets are constructed of sturdy 18-gauge (1-mm) thick double-wall, welded steel with 1-1/2” (38-mm) of insulating air space for fire resistance. High-performance, manual-close, self-latching doors open to a full 180 degrees and close easily and securely for maximum protection under fire conditions when pushed shut. Fail-safe closing mechanism ensures three-point stainless steel bullet latching system works every time. For added security, U-Loc™ handle comes with a cylinder lock and key set or accepts an optional padlock. Haz-Alert™ reflective warning labels are highly visible under fire conditions or during power outages when illuminated with a flashlight.Shelve adjust on 3” (76-mm) centers for versatile storage. Cabinets also feature a leakproof 2” (51-mm) bottom sump, four adjustable self-leveling feet, grounding connector and trilingual warning label. Durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance.All cabinets come with a ten-year warranty. Read More | [19 empty] | 23.0 | 8 | 22 | 17 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29070/name/1_Door%252C_Manual_Close%252C_Mini_Transportable_Flammable_Cabinet_for_Aerosols%252C_Sure-Grip%25C2%25AE_EX%252C_Yellow_-_890200.pdf/ |  |  | https://www.justrite.com/890200-sure-grip-ex-mini-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/8/9/890200_closed_lock.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/8/9/890200_closed_lock.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Piggyback 12 Gal (Manual) | 891300 | $964.00 | The Justrite Sure-Grip® EX Piggyback Safety Cabinet with Manual Close Doors (model 891300) is designed for the storage and segregation of flammable liquids in lab, construction, or engineering settings. With exterior dimensions of 18 in H x 43 in W x 18 in D (457 mm H x 1092 mm W x 457 mm D), its low-height design allows you to expand available storage by 12 gal (45 L) without taking up valuable floor space. Use independently on bench tops or stack on top of 30- or 45-gal (114- or 170-L) Justrite safety cabinets. It can also be mounted on a wall with an optional wall hanger assembly (model 25950). This yellow model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891300 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Piggyback Flammable Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 99.2 | 18 | 18 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29103/name/12_Gallon%252C_2_Doors%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Piggyback%252C_Yellow_-_891300.pdf/ |  |  | https://www.justrite.com/891300-sure-grip-ex-piggyback-flammable-safety-cabinet-yellow |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891300_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891300_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Piggyback 12 Gal (Self-Close, Yellow) | 891320 | $1,109.00 | The Justrite Sure-Grip® EX Piggyback Safety Cabinet with Self Close Doors (model 891320) is designed for the storage and segregation of flammable liquids in lab, construction, or engineering settings. With exterior dimensions of 18 in H x 43 in W x 18 in D (457 mm H x 1092 mm W x 457 mm D, its low-height design allows you to expand available storage by 12 gal (45 L) without taking up valuable floor space. Use independently on bench tops or stack on top of 30- or 45-gal (114- or 170-L) Justrite safety cabinets. It can also be mounted on a wall with an optional wall hanger assembly (model 25950). This yellow model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, self-closing doors open to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tight sequence. Continuous piano hinges ensure the doors close smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891320 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements. Safety Considerations for Piggyback Flammable Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more. Read More | [19 empty] | 109.0 | 18 | 18 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29107/name/12_Gallon%252C_2_Doors%252C_Self_Close%252C_Flammable_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Piggyback%252C_Yellow_-_891320.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-piggyback-flammable-safety-cabinet-cap-12-gallons-2-self-close-doors-yellow-891320 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891320_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891320_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Piggyback 12 Gallon (Self-Close) | 891325 | $1,253.00 | The Justrite Sure-Grip® EX Piggyback Safety Cabinet with Self Close Doors (model 891325) is designed for the storage and segregation of flammable liquids in lab, construction, or engineering settings. With exterior dimensions of 18 in H x 43 in W x 18 in D (457 mm H x 1092 mm W x 457 mm D, its low-height design allows you to expand available storage by 12 gal (45 L) without taking up valuable floor space. Use independently on bench tops or stack on top of 30- or 45-gal (114- or 170-L) Justrite safety cabinets. It can also be mounted on a wall with an optional wall hanger assembly (model 25950). This white model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, self-closing doors open to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tight sequence. Continuous piano hinges ensure the doors close smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891325 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements. Safety Considerations for Piggyback Flammable Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 110.0 | 18 | 18 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29110/name/12_Gallon%252C_2_Doors%252C_Self_Close%252C_Flammable_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Piggyback%252C_White_-_891325.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-piggyback-flammable-safety-cabinet-cap-12-gallons-2-self-close-doors-white-891325 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891325_3.png?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891325_3.png?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Piggyback 17 Gal (Manual) | 891700 | $1,111.00 | The Justrite Sure-Grip® EX Piggyback Safety Cabinet with Manual Close Doors (model 891700) is designed for the storage and segregation of flammable liquids in lab, construction, or engineering settings. With exterior dimensions of 24 in H x 43 in W x 18 in D (610 mm H x 1092 mm W x 457 mm D), its low-height design allows you to expand available storage by 17 gal (64 l) without taking up valuable floor space. Use independently on bench tops or stack on top of 30 or 45 gal (114/170 l) Justrite safety cabinets. It can also be mounted on a wall with an optional wall hanger assembly (model 25950). Looking for a custom safety cabinet? Request a quote here for your storage requirements. This yellow model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 891700 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29937. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-l), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891700 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Piggyback Flammable Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 123.7 | 18 | 24 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29111/name/17_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_Flammable_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Piggyback%252C_Yellow_-_891700.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-piggyback-flammable-safety-cabinet-cap-17-gallons-1-shelf-2-m-c-doors-yellow-891700 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891700-md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891700-md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Piggyback 17 Gal (Self-Close) | 891720 | $1,270.00 | The Justrite Sure-Grip® EX Piggyback Safety Cabinet with Self Close Doors (model 891720) is designed for the storage and segregation of flammable liquids in lab, construction, or engineering settings. With exterior dimensions of 24 in H x 43 in W x 18 in D (610 mm H x 1092 mm W x 457 mm D, its low-height design allows you to expand available storage by 17 gal (64 l) without taking up valuable floor space. Use independently on bench tops or stack on top of 30 or 45 gal (114/170 l) Justrite safety cabinets. It can also be mounted on a wall with an optional wall hanger assembly (model 25950). This yellow model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, self-closing doors open to a full 180 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tightly sequence. Continuous piano hinges ensure the doors close smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 891720 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29937. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-l), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891720 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements. Safety Considerations for Piggyback Flammable Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 125.5 | 18 | 24 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29115/name/17_Gallon%252C_1_Shelf%252C_2_Doors%252C_Self_Close%252C_Flammable_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Piggyback%252C_Yellow_-_891720.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-piggyback-flammable-safety-cabinet-cap-17-gallons-1-shelf-2-s-c-doors-yellow-891720 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891720-sfe-md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891720-sfe-md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Piggyback 17 Gal White (Self-Close) | 891725 | $1,389.00 | The Justrite Sure-Grip® EX Piggyback Safety Cabinet with Self Close Doors (model 891725) is designed for the storage and segregation of flammable liquids in lab, construction, or engineering settings. With exterior dimensions of 24 in H x 43 in W x 18 in D (610 mm H x 1092 mm W x 457 mm D, its low-height design allows you to expand available storage by 17 gal (64 L) without taking up valuable floor space. Use independently on bench tops or stack on top of 30- or 45-gal (114- or 170-L) Justrite safety cabinets. It can also be mounted on a wall with an optional wall hanger assembly (model 25950). This white model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, self-closing doors open to 140 degrees. A patented, concealed self-close mechanism provides obstruction-free access to shelf space. The doors self-index to guarantee they will close in tight sequence. Continuous piano hinges ensure the doors close smoothly every time.When used properly as designed, safety cabinets with self-closing doors shut automatically in case of a fire. The self-closing doors, under tension, are held open by using fusible links. Under fire conditions, when the ambient temperature reaches 165 F (74 C), the links melt and the safety cabinet’s self-closing doors automatically swing shut. The self-latching system engages with three stainless steel bullet-slams located at the top, bottom, and center locations of the door. For replacement fusible links, order model 29988. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 891725 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29937. The cabinet also features dual vents with flame arresters, four adjustable self-leveling feet, and a built-in grounding connector. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. All Justrite cabinets carry a 10-year warranty. Model 891725 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, NFPA 1, and IFC requirements. Safety Considerations for Piggyback Flammable Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Not just any safety cabinet will do. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article about the requirements for flammable liquid storage. Read More | [19 empty] | 128.0 | 18 | 24 | 43 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29118/name/17_Gallon%252C_1_Shelf%252C_2_Doors%252C_Self_Close%252C_Flammable_Safety_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Piggyback%252C_White_-_891725.pdf/ |  |  | https://www.justrite.com/sure-grip-ex-piggyback-flammable-safety-cabinet-cap-17-gallons-1-shelf-2-s-c-doors-white-891725 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_891725_white-safety-cabinet-main_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_891725_white-safety-cabinet-main_1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Slimline 22 Gal (Manual) | 892200 | $1,228.00 | Protect workers, reduce fire risks, and improve productivity by storing flammable fuels and chemicals in code-compliant safety cabinets. Designed to meet OSHA and NFPA standards, Sure-Grip® EX cabinets are constructed of sturdy 18-gauge (1-mm) thick double-wall, welded steel with 1-1/2" (38-mm) of insulating air space for fire resistance. High-performance, self-latching doors close easily and securely for maximum protection under fire conditions. Fail-safe closing mechanism ensures three-point stainless steel bullet latching system works every time. For added security, U-Loc™ handle comes with a cylinder lock and key set or accepts an optional padlock. Haz-Alert™ reflective warning labels are highly visible under fire conditions or during power outages when illuminated with a flashlight. Looking for a custom safety cabinet? Request a quote here for your storage requirements. Patented SpillSlope® galvanized steel shelves direct spills to back and bottom of a leakproof 2” (51-mm) bottom sump. They adjust on 3” (76-mm) centers for versatile storage. Cabinets also feature dual vents with flame arresters, four adjustable self-leveling feet, grounding connector, and trilingual warning label. Durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance. Manual-close door(s) opens to a full 180 degrees and self-latches when pushed closed. Self-close door(s) shuts and latches automatically when a fusible link melts at 165°F (74ºC) under fire conditions. Unique, concealed self-closing mechanism offers obstruction-free access to contents. All door styles meet OSHA and NFPA 30; self-close door styles also meet NFPA 1 and the International Fire Code. Most cabinets are FM approved and all come with a ten-year warranty.Use alone or alongside larger cabinets for expanded storage when floor space is limited. The standard slimline cabinet is 18” (457-mm) deep and fits flush alongside a 45-gallon (170-L) cabinet. The deep slimline cabinet is 34” (864-mm) deep and fits alongside a 60- or 90-gallon (227- or 341-L) cabinet. Both cabinets measure only 23-1/4” (591-mm) wide. These cabinets have all the great features and benefits of our classic cabinets as outlined below. Read More | [19 empty] | 189.0 | 18 | 65 | 23.25 |  |  |  | https://www.justrite.com/productpdf/download/file/id/30386/name/22_gallon_Flammable_Safety_Cabinet%252C_Slimline%252C_1_Manual_Close_Door_-_Sure-Grip%25C2%25AE_EX.pdf/ |  |  | https://www.justrite.com/sure-gripr-ex-slimline-flammable-safety-cabinet-cap-22-gallons-3-shelves-1-m-c-door |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_9_892200-md_1_4.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_9_892200-md_1_4.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Under Fume Hood (15 Gal.) | 882400 | $1,641.00 | Keeping bottles of chemicals in a fume hood is a safety risk. The Justrite Sure-Grip® EX Under Fume Hood Safety Cabinet with 1 Manual Close Door, Right Hinge (model 882400) is designed for the safe and code-compliant storage and venting of solvents and flammable liquids in laboratory environments within colleges, universities, pharmaceutical R&D settings, and more.The yellow cabinet can store 15 gal (57 L) of flammable liquids and fits under chemical fume hoods, benches, and counters. It features an ergonomic design with a 4-in high by 3-in deep (102-mm high by 77-mm deep) toe kick, which makes it easier for chemical handlers to maintain an appropriate distance. Its exterior dimensions are 35.75 in H x 24 in W x 21.625 in D (908 mm H x 610 mm W x 549 mm D); its 35.75-in (908-mm) height includes leg levelers installed and extended 5/8-in (15.9-mm). This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, manual-close, self-latching door (fail safe, 3-point stainless steel) with a right hinge closure can open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the door closes smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 882400 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29950.The cabinet also features back wall vents with built-in flame arresters and a grounding connector to electrically bond directly to the chemical fume hood. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. Our line of under fume hood flammable cabinets are available in four widths ranging from 24 in (610 mm) to 48 in (1219 mm). Single-door cabinets are offered in either left or right-hinge options for easy operation.All Justrite cabinets carry a 10-year warranty. Model 882400 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Under Fume Hood Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article, about the requirements for flammable liquid storage. Read More | [19 empty] | 115.0 | 21.625 | 35.75 | 24 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29143/name/15_Gallon%252C_1_Shelf%252C_1_Door%252C_Manual_Close%252C_Right_Hinge%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Under_Fume_Hood%252C_Yellow_-_882400.pdf/ |  |  | https://www.justrite.com/under-fume-hood-solvent-flammable-liquid-safety-cabinet-cap-15-gal-1-shlf-1-m-c-rgt-door-yel-882400 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_8_882400_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_8_882400_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Under Fume Hood Safety Cab. (19 Gal.) | 883000 | $1,785.00 | Keeping bottles of chemicals in a fume hood is a safety risk. The Justrite Sure-Grip® EX Under Fume Hood Safety Cabinet with Manual Close Doors (model 883000) is designed for the safe and code-compliant storage and venting of solvents and flammable liquids in laboratory environments within colleges, universities, pharmaceutical R&D settings, and more.The yellow cabinet can store 19 gal (72 L) of flammable liquids and fits under chemical fume hoods, benches, and counters. It features an ergonomic design with a 4-in high by 3-in deep (102-mm high by 77-mm deep) toe kick, which makes it easier for chemical handlers to maintain an appropriate distance. Its exterior dimensions are 35.75 in H x 30 in W x 21.625 in D (908 mm H x 762 mm W x 549 mm D); its 35.75-in (908-mm) height includes leg levelers installed and extended 5/8-in (15.9-mm). This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. High-performance, manual-close, self-latching doors (fail safe, 3-point stainless steel) open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the doors close smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 883000 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29951.The cabinet also features back wall vents with built-in flame arresters and a grounding connector to electrically bond directly to the chemical fume hood. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. Our line of under fume hood flammable cabinets are available in four widths ranging from 24 in (610 mm) to 48 in (1219 mm). Single-door cabinets are offered in either left or right-hinge options for easy operation.All Justrite cabinets carry a 10-year warranty. Model 883000 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Under Fume Hood Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article, about the requirements for flammable liquid storage. Read More | [19 empty] | 156.0 | 21.625 | 35.75 | 30 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29155/name/19_Gallon%252C_1_Shelf%252C_2_Doors%252C_Manual_Close%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Under_Fume_Hood%252C_Yellow_-_883000.pdf/ |  |  | https://www.justrite.com/under-fume-hood-solvent-flammable-liquid-safety-cabinet-cap-19-gal-1-shelf-2-m-c-doors-yellow-883000 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_8_883000_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_8_883000_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Sure-Grip EX Under Fume Hood Solvent/Flammable 15 | 882410 | $1,683.00 | Keeping bottles of chemicals in a fume hood is a safety risk. The Justrite Sure-Grip® EX Under Fume Hood Safety Cabinet with 1 Manual Close Door, Left Hinge (model 882410) is designed for the safe and code-compliant storage and venting of solvents and flammable liquids in laboratory environments within colleges, universities, pharmaceutical R&D settings, and more.The yellow cabinet can store 15 gal (57 L) of flammable liquids and fits under chemical fume hoods, benches, and counters. It features an ergonomic design with a 4-in high by 3-in deep (102-mm high by 77-mm deep) toe kick, which makes it easier for chemical handlers to maintain an appropriate distance. Its exterior dimensions are 35.75 in H x 24 in W x 21.625 in D (908 mm H x 610 mm W x 549 mm D); its 35.75-in (908-mm) height includes leg levelers installed and extended 5/8-in (15.9-mm). This model is constructed with a sturdy 18-gauge (1-mm) thick steel double-wall, with 1.5 in (38 mm) of insulating air space for fire resistance. The fully welded, not riveted, construction holds the squareness of the unit for a longer life and offers greater protection in a fire since air gaps are reduced. The high-performance, manual-close, self-latching door (fail safe, 3-point stainless steel) with a left hinge closure can open to a full 180 degrees and close securely for maximum protection under fire conditions when pushed shut. Continuous piano hinges ensure the door closes smoothly every time. For added security, each cabinet comes with an innovative U-Loc™ handle that does not require manual rotation to engage the mandatory latch. Cabinets include a cylinder lock and key set, or you can choose to use your own padlock. Haz-Alert™ reflective warning labels that read “FLAMMABLE – KEEP FIRE AWAY” are highly visible under fire conditions or during power outages when illuminated with a flashlight.Model 882410 features one patented, galvanized steel SpillSlope® shelf designed to direct spills to the back and bottom into a 2-in (51-mm) leak-tested sump that meets EPA requirements. Adjustable on 3-in (76-mm) centers, all SpillSlope shelves meet ANSI MH28.1 standards with a 350-lb (159-kg) safe allowable load. Welded shelf hangers interlock with the shelf to offer maximum stability with no flimsy brackets to lose. For replacement shelves, use model 29950.The cabinet also features back wall vents with built-in flame arresters and a grounding connector to electrically bond directly to the chemical fume hood. Its durable lead-free epoxy/polyester powder-coat finish provides excellent chemical resistance inside and out, as well as minimizes the effects of corrosion and humidity.Justrite supplies more than 300 safety cabinet models ranging from 4- to 115-gal (15- to 435-L), to meet a versatile range of safety, compliance, and storage requirements. Our line of under fume hood flammable cabinets are available in four widths ranging from 24 in (610 mm) to 48 in (1219 mm). Single-door cabinets are offered in either left or right-hinge options for easy operation.All Justrite cabinets carry a 10-year warranty. Model 882410 is tested and approved by FM Approvals. It complies with OSHA, NFPA 30, and NFPA 1 requirements. Safety Considerations for Under Fume Hood Cabinets The requirements for handling and storing chemicals are complex and depend on how flammable the liquid is. There is no room for error, so it is critical to choose a safety cabinet that has been independently tested to meet the standards of your industry and local jurisdictions. Justrite flammable cabinets are best in class and are trusted within a wide range of industrial settings across the world, including laboratories, manufacturing, oil refineries, chemical processing plants, and more.For more information, please consult our workplace safety article, about the requirements for flammable liquid storage. Read More | [19 empty] | 115.0 | 21.625 | 35.75 | 24 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29146/name/15_Gallon%252C1_Shelf%252C_1_Door%252C_Manual_Close%252C_Left_Hinge%252C_Flammable_Cabinet%252C_Sure-Grip%25C2%25AE_EX_Under_Fume_Hood%252C_Yellow_-_882410.pdf/ |  |  | https://www.justrite.com/under-fume-hood-solvent-flammable-liquid-safety-cabinet-cap-15-gal-1-shlf-1-m-c-lft-door-yel-882410 |  |  |  | https://www.justrite.com/media/catalog/product/_/8/_8_8_882410_md_3.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/8/_8_8_882410_md_3.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | 2 Drum DrumShed With Rolltop Doors, EcoPolyBlend | 28675 | $3,531.00 | EcoPolyBlend™ DrumSheds™ feature a weather resistant, robust design made of eco-friendly 50% recycled polyethylene. Convenient roll top doors allow access from front or rear. Wide door opening of 52" accepts palletized loads. The leak proof pallet sump meets EPA and SPCC requirements and has removable, durable deck grating for easy clean up of spills. A spacious interior offers ample headroom for adding funnels or pumps. Safety features include grounding capability, integral anchor bolt locations for wind resistance, and molded in lugs which accept padlocks for security. Forliftk pockets front and back make it easy to relocate empty sheds. Ships fully assembled.Optional pallet ramp for easy loading.EPA and SPCC Compliant(UDL) Uniformly Distributed Load - 1250 Lbs/567 Kgs. per drum, for DrumSheds™ and Pallets | [19 empty] | 285.3 | 58.5 | 75.25 | 60.75 |  |  |  | https://www.justrite.com/productpdf/download/file/id/29855/name/2_Drum_DrumShed%25E2%2584%25A2_With_Rolltop_Doors%252C_EcoPolyBlend%25E2%2584%25A2_-_28675.pdf/ |  |  | https://www.justrite.com/ecopolyblend-drumshed-with-rolltop-doors-accommodates-2-drums-optional-ramp-poly-28675 |  |  |  | https://www.justrite.com/media/catalog/product/_/2/_2_8_28675-lg_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/2/_2_8_28675-lg_1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Steel Gas Cylinder Stand 35282 | 35282 | $240.00 | Designed for the safe storage of industrial and commercial use compressed gas cylinders, these storage stands help prevent dangerous tipping situations. Powder-coated to ensure long service life and chemical resistance. Features a 0.187-in (5-mm) thick steel clamping bar protected with an edge bumper guard that attaches to the stand using dual threaded rods and wing nuts. Pre-drilled holes in base for floor mounting with application specific fasteners (not included).Heavy-duty polypropylene support strap, 1.5- x 54-in (38-mm x 1372-mm) and rated at 1138-lb (516-kg), features steel cinch buckle and hold individual cylinders securely against stand. Helps comply with OSHA, NFP1, NFPA 55, and CGA. | [19 empty] | 21 | 16 | 15 | 16 |  |  |  | https://www.justrite.com/productpdf/download/file/id/30870/name/Steel_Gas_Cylinder_Stand%252C_1_Cylinder_Capacity%252C_Low_Profile_-_35282.pdf/ |  |  | https://www.justrite.com/gas-cylinder-stand-1-cylinder-capacity-low-profile-steel-35282 |  |  |  | https://www.justrite.com/media/catalog/product/_/3/_3_5_35282-lg.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/3/_3_5_35282-lg.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | HDPE Gas Cylinder Support Bracket With Chain | 35256 | $240.00 | Designed for the safe storage of industrial and commercial use compressed gas cylinders, these storage stands help prevent dangerous tipping situations. Powder-coated to ensure long service life and chemical resistance. Features a 0.187-in (5-mm) thick steel clamping bar protected with an edge bumper guard that attaches to the stand using dual threaded rods and wing nuts. Pre-drilled holes in base for floor mounting with application specific fasteners (not included).Heavy-duty polypropylene support strap, 1.5- x 54-in (38-mm x 1372-mm) and rated at 1138-lb (516-kg), features steel cinch buckle and hold individual cylinders securely against stand. Helps comply with OSHA, NFP1, NFPA 55, and CGA. | [19 empty] | 21 | 16 | 15 | 16 |  |  |  | https://www.justrite.com/productpdf/download/file/id/30870/name/Steel_Gas_Cylinder_Stand%252C_1_Cylinder_Capacity%252C_Low_Profile_-_35282.pdf/ |  |  | https://www.justrite.com/gas-cylinder-stand-1-cylinder-capacity-low-profile-steel-35282 |  |  |  | https://www.justrite.com/media/catalog/product/_/3/_3_5_35282-lg.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/3/_3_5_35282-lg.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Oily Waste Can (10 Gal) | 9300 | $101.00 | Rags and cloths soaked with solvents, thinners, linseed oil, combustible adhesives and other flammable liquids present a serious fire risk when improperly discarded. Lid opens to no more than 60 degrees and stays closed when not in use, isolating contents from fire sources and limiting oxygen to virtually eliminate the risk of spontaneous combustion. Round construction and elevated base allow circulation of air around can to disperse heat and reduce moisture build-up and rusting. Durable steel construction with a powder-coat finish provides chemical resistance. Ten-year warranty. | [19 empty] | 10.5 |  | 18.25 |  |  |  |  | https://www.justrite.com/productpdf/download/file/id/30289/name/10_Gallon%252C_Oily_Waste_Can%252C_Hands-Free%252C_Self-Closing_Cover%252C_Red_-_09300.pdf/ |  |  | https://www.justrite.com/09300-oily-waste-can-10-gallon-foot-operated-self-closing-cover-red |  |  |  | https://www.justrite.com/media/catalog/product/_/0/_0_9_09300_oily-waste-can-10-gallon-red_justrite_1_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/0/_0_9_09300_oily-waste-can-10-gallon-red_justrite_1_1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Oily Waste Can (14 Gal) | 9500 | $116.00 | Rags and cloths soaked with solvents, thinners, linseed oil, combustible adhesives and other flammable liquids present a serious fire risk when improperly discarded. Lid opens to no more than 60 degrees and stays closed when not in use, isolating contents from fire sources and limiting oxygen to virtually eliminate the risk of spontaneous combustion. Round construction and elevated base allow circulation of air around can to disperse heat and reduce moisture build-up and rusting. Durable steel construction with a powder-coat finish provides chemical resistance. Ten-year warranty. | [19 empty] | 14.0 |  | 20.25 |  |  |  |  | https://www.justrite.com/productpdf/download/file/id/30297/name/14_Gallon%252C_Oily_Waste_Can%252C_Hands-Free%252C_Self-Closing_Cover%252C_Red_-_09500.pdf/ |  |  | https://www.justrite.com/09500-oily-waste-can-14-gallon-foot-operated-self-closing-cover-red |  |  |  | https://www.justrite.com/media/catalog/product/_/0/_0_9_09500_oily-waste-can-14-gallon-red_justrite_1_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/0/_0_9_09500_oily-waste-can-14-gallon-red_justrite_1_1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Oily Waste Can (21 Gal) | 9700 | $172.00 | Rags and cloths soaked with solvents, thinners, linseed oil, combustible adhesives and other flammable liquids present a serious fire risk when improperly discarded. Lid opens to no more than 60 degrees and stays closed when not in use, isolating contents from fire sources and limiting oxygen to virtually eliminate the risk of spontaneous combustion. Round construction and elevated base allow circulation of air around can to disperse heat and reduce moisture build-up and rusting. Durable steel construction with a powder-coat finish provides chemical resistance. Ten-year warranty. | [19 empty] | 18.8 |  | 23.438 |  |  |  |  | https://www.justrite.com/productpdf/download/file/id/30303/name/21_Gallon%252C_Oily_Waste_Can%252C_Hands-Free%252C_Self-Closing_Cover%252C_Red_-_09700.pdf/ |  |  | https://www.justrite.com/oily-waste-can-21-gallon-foot-operated-self-closing-cover-red |  |  |  | https://www.justrite.com/media/catalog/product/_/0/_0_9_09700_oily-waste-can-21-gallon-red_justrite_1_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/0/_0_9_09700_oily-waste-can-21-gallon-red_justrite_1_1.jpg?store=justrite_store_view&image-type=image
https://www.justrite.com/ | Justrite Manufacturing | Oily Waste Can (6 Gal) | 9100 | $88.00 | Rags and cloths soaked with solvents, thinners, linseed oil, combustible adhesives and other flammable liquids present a serious fire risk when improperly discarded. Lid opens to no more than 60 degrees and stays closed when not in use, isolating contents from fire sources and limiting oxygen to virtually eliminate the risk of spontaneous combustion. Round construction and elevated base allow circulation of air around can to disperse heat and reduce moisture build-up and rusting. Durable steel construction with a powder-coat finish provides chemical resistance.Ten-year warranty. | [19 empty] | 8.3 |  | 15.875 |  |  |  |  | https://www.justrite.com/productpdf/download/file/id/30282/name/6_Gallon%252C_Steel_Oily_Waste_Can%252C_Hands-Free%252C_Self-Closing_Cover%252C_Red_-_09100.pdf/ |  |  | https://www.justrite.com/09100-oily-waste-can-6-gallon-foot-operated-self-closing-cover-red |  |  |  | https://www.justrite.com/media/catalog/product/_/0/_0_9_09100_1.jpg?store=justrite_store_view&image-type=image | https://www.justrite.com/media/catalog/product/_/0/_0_9_09100_1.jpg?store=justrite_store_view&image-type=image